## QAReport
Hospital: |  |  |  |  | Model: |  | Model
Department / Room: |  | Reporting Monitor X |  |  | System no.:
Installation Date: |  | 2025-01-01 |  |  | Software version: |  | VXXX
Date of visit: |  | 2025-01-01 |  |  | Manufacturer: |  | Manufacturer
Date of report: |  | 2025-01-01 |  |  | Serial number:
Survey & Report by:
COMMENTS
All results are within expected tolerances.
The majority of the results are within expected tolerances.
CONCLUSIONS
The monitors are suitable for clinical use.
SUMMARY OF TESTS & RESULTS | [9 empty] | Expected values
1. Monitors |  |  |  | Results |  | Pass/Fail | Comments | P | O
1.1 | Ambient Light |  |  | 4.5 |  | O |  | 25-75 | AAPM TG270 jan-19, table X
1.2 | Lmax difference |  |  | 4.113% |  | P |  | <10
1.3 | Lamb |  |  | 0 |  | P |  | 0.13375 | AAPM TG270 jan-19, table X |  |  |  |  | Reporting rm
| Right Monitor | [10 empty] | Calibration |  | DICOM
1.4 | Minimum Luminance |  |  | 0.535 |  | P |  | <2 |  |  |  | Last calibration |  | 2023-11-03
1.5 | Maximum Luminance |  |  | 462 |  | P |  | >450 | AAPM TG270 jan-19, table X |  |  | Use external sensor (Eizo sensor) |  | P
1.6 | Luminance ratio |  |  | 863.5514 |  |  |  | 850 |  |  |  | target luminance:
1.7 | Effective Luminance ratio |  |  | 863.5514 |  | P |  | >250 |  |  |  |  | - Lmax (cd/m2) | 500
1.8 | Eff Luminance/Luminance |  |  | 1 |  | P |  | >0.6 |  |  |  |  | - Lmin (cd/m2) | 0.6
1.9 | DICOM GSDF calibration |  |  | 3.791% |  | P |  | <10 |  |  |  | Apply ambient light:
1.1 | Non-uniformity |  | UN10 | 9.959% |  | P |  | >30 |  |  |  |  | - Ambient light (lux) | 0
|  |  | UN80 | 3.636% |  | P |  | >30 |  |  |  | Scheduled tests:
1.11 | Artefacts |  |  | none present | [6 empty] | 1.92 |  | - consistency tests | monthly
| Left Monitor | [11 empty] | - calibration | annually
1.12 | Minimum Luminance |  |  | 0.601 |  | P |  | <2
1.13 | Maximum Luminance |  |  | 481 |  | P |  | >450 | AAPM TG270 jan-19, table X
1.14 | Luminance ratio |  |  | 800.3328 |  |  |  | 850
1.15 | Effective Luminance ratio |  |  | 800.3328 |  | P |  | >250
1.16 | Eff Luminance/Luminance |  |  | 1 |  | P |  | >0.6
1.17 | DICOM GSDF calibration |  |  | 7.173% |  | P |  | <10
1.18 | Non-uniformity |  | UN10 | 7.646% |  | P |  | >30
|  |  | UN80 | 2.02% |  | P |  | >30
1.19 | Artefacts |  |  | none present |  |  |  | <10
Note: Tolerances have been updated to adhere to AAPM TG270 report (Jan 2019), a summary of which is found in table X (below). In particular, ambient light (Lamb) is now required to be below the minimum value between the RHS and LHS monitors of 0.25 x minimum luminance (Lmin, with no correction for ambient light).
TG270 Report:
Acceptance Report

## TG18
, Reporting Monitor X
2025-01-01
Ambient Light | [9 empty] | TG18-QC | TG18-LN
Illuminance (lux) |  | 4.5
Luminance Response
Left Monitor
Ambient Luminance Lamb (cd/m²) |  |  | 0 |  | Bit depth | 12
DICOM 3.14 standard:
test pattern |  | Luminance (cd/m2) |  | JND | dL/L | dL/L for JND | average | diff from |  | test pattern |  | Luminance | JND | dL/L | dL/L |  | JND |  | dL/L for JND |  |  | average
TG18-LN | p-value | Excl amb | Incl amb | j(L) |  |  | JND | DICOM (%) |  | TG18 | p-value | L | j(L) |  | +10% | -10% | +10% | -10% | 0% | +10% | -10% | JND
1 | 0 | 0.601 | 0.601 | 52.5255 | [5 empty] | 1 | 0 | 0.601 | 52.5255 |  |  |  | 57.7781 | 47.273
2 | 240 | 1.51 | 1.51 | 90.6219 | 0.8612 | 0.022606 | 71.5737 | 1.4926 |  | 2 | 240 | 1.5345 | 90.6219 | 0.87425 | 0.96168 | 0.78683 | 99.6841 | 81.5597 | 0.022948 | 0.025243 | 0.020654 | 71.5737
3 | 480 | 3.05 | 3.05 | 128.7183 | 0.67544 | 0.01773 | 109.6701 | 1.1584 |  | 3 | 480 | 3.0725 | 128.7183 | 0.6677 | 0.73447 | 0.60093 | 141.5901 | 115.8465 | 0.017527 | 0.019279 | 0.015774 | 109.6701
4 | 720 | 5.43 | 5.43 | 166.8147 | 0.56132 | 0.014734 | 147.7665 | 2.4513 |  | 4 | 720 | 5.3911 | 166.8147 | 0.54789 | 0.60268 | 0.4931 | 183.4961 | 150.1332 | 0.014382 | 0.01582 | 0.012944 | 147.7665
5 | 960 | 8.72 | 8.72 | 204.911 | 0.46502 | 0.012206 | 185.8629 | 1.2423 |  | 5 | 960 | 8.7113 | 204.911 | 0.47087 | 0.51795 | 0.42378 | 225.4021 | 184.4199 | 0.01236 | 0.013596 | 0.011124 | 185.8629
6 | 1200 | 13.2 | 13.2 | 243.0074 | 0.40876 | 0.01073 | 223.9592 | 2.2 |  | 6 | 1200 | 13.314 | 243.0074 | 0.41795 | 0.45975 | 0.37616 | 267.3082 | 218.7067 | 0.010971 | 0.012068 | 0.0098739 | 223.9592
7 | 1440 | 19.5 | 19.5 | 281.1038 | 0.38532 | 0.010114 | 262.0556 | 1.4373 |  | 7 | 1440 | 19.5573 | 281.1038 | 0.37986 | 0.41785 | 0.34188 | 309.2142 | 252.9934 | 0.0099711 | 0.010968 | 0.008974 | 262.0556
8 | 1680 | 27.9 | 27.9 | 319.2002 | 0.35443 | 0.0093035 | 300.152 | 0.8396 |  | 8 | 1680 | 27.8969 | 319.2002 | 0.35148 | 0.38663 | 0.31633 | 351.1202 | 287.2802 | 0.0092261 | 0.010149 | 0.0083035 | 300.152
9 | 1920 | 38.9 | 38.9 | 357.2966 | 0.32934 | 0.008645 | 338.2484 | 0.13236 |  | 9 | 1920 | 38.9131 | 357.2966 | 0.32978 | 0.36276 | 0.2968 | 393.0262 | 321.5669 | 0.0086564 | 0.009522 | 0.0077908 | 338.2484
10 | 2160 | 53.5 | 53.5 | 395.3929 | 0.31602 | 0.0082952 | 376.3447 | 1.0116 |  | 10 | 2160 | 53.3447 | 395.3929 | 0.31285 | 0.34414 | 0.28157 | 434.9322 | 355.8536 | 0.0082121 | 0.0090334 | 0.0073909 | 376.3447
11 | 2400 | 72.3 | 72.3 | 433.4893 | 0.29889 | 0.0078456 | 414.4411 | 0.18816 |  | 11 | 2400 | 72.1317 | 433.4893 | 0.29945 | 0.3294 | 0.26951 | 476.8382 | 390.1404 | 0.0078603 | 0.0086464 | 0.0070743 | 414.4411
12 | 2640 | 97 | 97 | 471.5857 | 0.29179 | 0.0076593 | 452.5375 | 1.0651 |  | 12 | 2640 | 96.4707 | 471.5857 | 0.28871 | 0.31759 | 0.25984 | 518.7443 | 424.4271 | 0.0075785 | 0.0083364 | 0.0068207 | 452.5375
13 | 2880 | 129 | 129 | 509.6821 | 0.28319 | 0.0074334 | 490.6339 | 1.123 |  | 13 | 2880 | 127.8851 | 509.6821 | 0.28004 | 0.30804 | 0.25204 | 560.6503 | 458.7139 | 0.0073509 | 0.0080859 | 0.0066158 | 490.6339
14 | 3120 | 170 | 170 | 547.7784 | 0.27425 | 0.0071988 | 528.7303 | 0.4603 |  | 14 | 3120 | 168.315 | 547.7784 | 0.27299 | 0.30029 | 0.24569 | 602.5563 | 493.0006 | 0.0071658 | 0.0078824 | 0.0064492 | 528.7303
15 | 3360 | 223 | 223 | 585.8748 | 0.26972 | 0.0070799 | 566.8266 | 0.92789 |  | 15 | 3360 | 220.2329 | 585.8748 | 0.26724 | 0.29396 | 0.24052 | 644.4623 | 527.2873 | 0.0070149 | 0.0077163 | 0.0063134 | 566.8266
16 | 3600 | 291 | 291 | 623.9712 | 0.26459 | 0.0069453 | 604.923 | 0.77989 |  | 16 | 3600 | 286.7909 | 623.9712 | 0.26254 | 0.2888 | 0.23629 | 686.3683 | 561.5741 | 0.0068916 | 0.0075807 | 0.0062024 | 604.923
17 | 3840 | 379 | 379 | 662.0676 | 0.26269 | 0.0068953 | 643.0194 | 1.5361 |  | 17 | 3840 | 372.011 | 662.0676 | 0.25871 | 0.28458 | 0.23284 | 728.2743 | 595.8608 | 0.006791 | 0.0074701 | 0.0061119 | 643.0194
18 | 4080 | 481 | 481 | 700.164 | 0.23721 | 0.0062266 | 681.1158 | 7.1731 |  | 18 | 4080 | 481 | 700.164 | 0.25554 | 0.28109 | 0.22999 | 770.1804 | 630.1476 | 0.0067077 | 0.0073785 | 0.0060369 | 681.1158
Lmin |  | 0.601 |  | Luminance ratio | 800.3328
Lmax |  | 481 |  | Eff. Luminance ratio | 800.3328 | Eff Lum/Lum = |  | 1
linear relation between p-value and JND:
[11 empty] | a: | 0.15873
[11 empty] | b: | 52.5255
Right Monitor
Ambient Luminance (cd/m²) |  |  | 0 |  | Bit depth | 12
DICOM 3.14 standard:
test pattern |  | Luminance (cd/m2) |  | JND | dL/L | dL/L for JND | average | diff from |  | test pattern |  | Luminance | JND | dL/L |  |  | JND |  | dL/L for JND |  |  | average
TG18-LN | p-value | Excl amb | Incl amb | j(L) |  |  | JND | DICOM (%) |  | TG18 | p-value | L | j(L) | 0% | +10% | -10% | +10% | -10% | 0% | +10% | -10% | JND
1 | 0 | 0.535 | 0.535 | 48.7012 | [5 empty] | 1 | 0 | 0.535 | 48.7012 |  |  |  | 53.5714 | 43.8311 | dL/L | +10% | -10%
2 | 240 | 1.41 | 1.41 | 86.6697 | 0.89974 | 0.023697 | 67.6855 | 0.13888 |  | 2 | 240 | 1.4122 | 86.6697 | 0.90099 | 0.99109 | 0.81089 | 95.3367 | 78.0027 | 0.02373 | 0.026103 | 0.021357 | 67.6855
3 | 480 | 2.87 | 2.87 | 124.6382 | 0.68224 | 0.017969 | 105.6539 | 0.027076 |  | 3 | 480 | 2.8739 | 124.6382 | 0.68206 | 0.75026 | 0.61385 | 137.102 | 112.1743 | 0.017964 | 0.01976 | 0.016167 | 105.6539
4 | 720 | 5.14 | 5.14 | 162.6066 | 0.56679 | 0.014928 | 143.6224 | 1.8346 |  | 4 | 720 | 5.0903 | 162.6066 | 0.55658 | 0.61224 | 0.50092 | 178.8673 | 146.346 | 0.014659 | 0.016125 | 0.013193 | 143.6224
5 | 960 | 8.3 | 8.3 | 200.5751 | 0.47024 | 0.012385 | 181.5908 | 1.3182 |  | 5 | 960 | 8.2746 | 200.5751 | 0.47652 | 0.52417 | 0.42887 | 220.6326 | 180.5176 | 0.01255 | 0.013805 | 0.011295 | 181.5908
6 | 1200 | 12.6 | 12.6 | 238.5435 | 0.41148 | 0.010838 | 219.5593 | 2.4455 |  | 6 | 1200 | 12.6976 | 238.5435 | 0.4218 | 0.46398 | 0.37962 | 262.3979 | 214.6892 | 0.011109 | 0.01222 | 0.0099983 | 219.5593
7 | 1440 | 18.7 | 18.7 | 276.512 | 0.38978 | 0.010266 | 257.5278 | 1.8903 |  | 7 | 1440 | 18.7038 | 276.512 | 0.38255 | 0.4208 | 0.34429 | 304.1632 | 248.8608 | 0.010075 | 0.011083 | 0.0090678 | 257.5278
8 | 1680 | 26.6 | 26.6 | 314.4804 | 0.34879 | 0.0091862 | 295.4962 | 1.2985 |  | 8 | 1680 | 26.7316 | 314.4804 | 0.35337 | 0.38871 | 0.31804 | 345.9285 | 283.0324 | 0.0093071 | 0.010238 | 0.0083764 | 295.4962
9 | 1920 | 37 | 37 | 352.4489 | 0.32704 | 0.0086136 | 333.4647 | 1.2288 |  | 9 | 1920 | 37.3389 | 352.4489 | 0.33111 | 0.36422 | 0.298 | 387.6938 | 317.204 | 0.0087207 | 0.0095928 | 0.0078487 | 333.4647
10 | 2160 | 51.1 | 51.1 | 390.4174 | 0.32009 | 0.0084304 | 371.4331 | 2.013 |  | 10 | 2160 | 51.235 | 390.4174 | 0.31377 | 0.34515 | 0.2824 | 429.4591 | 351.3756 | 0.0082641 | 0.0090905 | 0.0074377 | 371.4331
11 | 2400 | 69.3 | 69.3 | 428.3858 | 0.30233 | 0.0079625 | 409.4016 | 0.75555 |  | 11 | 2400 | 69.3221 | 428.3858 | 0.30006 | 0.33006 | 0.27005 | 471.2244 | 385.5472 | 0.0079028 | 0.0086931 | 0.0071126 | 409.4016
12 | 2640 | 92.6 | 92.6 | 466.3543 | 0.28783 | 0.0075808 | 447.37 | 0.43098 |  | 12 | 2640 | 92.7475 | 466.3543 | 0.28908 | 0.31799 | 0.26017 | 512.9897 | 419.7189 | 0.0076136 | 0.008375 | 0.0068523 | 447.37
13 | 2880 | 123 | 123 | 504.3227 | 0.282 | 0.0074273 | 485.3385 | 0.64059 |  | 13 | 2880 | 122.9705 | 504.3227 | 0.28021 | 0.30823 | 0.25219 | 554.755 | 453.8905 | 0.00738 | 0.008118 | 0.006642 | 485.3385
14 | 3120 | 163 | 163 | 542.2912 | 0.27972 | 0.0073672 | 523.307 | 2.4616 |  | 14 | 3120 | 161.8482 | 542.2912 | 0.273 | 0.3003 | 0.2457 | 596.5203 | 488.0621 | 0.0071902 | 0.0079092 | 0.0064712 | 523.307
15 | 3360 | 215 | 215 | 580.2597 | 0.27513 | 0.0072463 | 561.2754 | 3.0002 |  | 15 | 3360 | 211.745 | 580.2597 | 0.26712 | 0.29383 | 0.24041 | 638.2856 | 522.2337 | 0.0070353 | 0.0077388 | 0.0063317 | 561.2754
16 | 3600 | 281 | 281 | 618.2281 | 0.26613 | 0.0070092 | 599.2439 | 1.4554 |  | 16 | 3600 | 275.6726 | 618.2281 | 0.26231 | 0.28854 | 0.23608 | 680.0509 | 556.4053 | 0.0069087 | 0.0075995 | 0.0062178 | 599.2439
17 | 3840 | 361 | 361 | 656.1966 | 0.24922 | 0.0065639 | 637.2123 | 3.5469 |  | 17 | 3840 | 357.4702 | 656.1966 | 0.25839 | 0.28422 | 0.23255 | 721.8162 | 590.5769 | 0.0068053 | 0.0074858 | 0.0061247 | 637.2123
18 | 4080 | 462 | 462 | 694.165 | 0.24544 | 0.0064644 | 675.1808 | 3.7913 |  | 18 | 4080 | 462 | 694.165 | 0.25512 | 0.28063 | 0.2296 | 763.5815 | 624.7485 | 0.0067191 | 0.0073911 | 0.0060472 | 675.1808
Lmin |  | 0.535 |  | Luminance ratio | 863.5514
Lmax |  | 462 |  | Eff. Luminance ratio | 863.5514 | Eff Lum/Lum = |  | 1
[9 empty] | , Reporting Monitor X |  |  | linear relation between p-value and JND:
[9 empty] | 2025-01-01 |  | a: | 0.1582
[11 empty] | b: | 48.7012
Luminance non-uniformity
Test Pattern TG18-UN10
Left monitor
Position |  | Centre | Top R | Top L | Bottom R | Bottom L
Luminance (cd/m²) |  | 2.42 | 2.39 | 2.58 | 2.39 | 2.56
Non-uniformity (%) |  |  | 1.2397 | -6.6116 | 1.2397 | -5.7851
max non-uniformity (%) |  | 7.6459
Right monitor
Position |  | Centre | Top R | Top L | Bottom R | Bottom L
Luminance (cd/m²) |  | 2.29 | 2.53 | 2.32 | 2.3 | 2.3
Non-uniformity (%) |  |  | -10.4803 | -1.31 | -0.43668 | -0.43668
max non-uniformity (%) |  | 9.9585
Test Pattern TG18-UN80
Left monitor
Position |  | Centre | Top R | Top L | Bottom R | Bottom L
Luminance (cd/m²) |  | 200 | 200 | 196 | 199 | 197
Non-uniformity (%) |  |  | 0 | 2 | 0.5 | 1.5
max non-uniformity (%) |  | 2.0202
Right monitor
Position |  | Centre | Top R | Top L | Bottom R | Bottom L
Luminance (cd/m²) |  | 190 | 196 | 192 | 195 | 189
Non-uniformity (%) |  |  | -3.1579 | -1.0526 | -2.6316 | 0.52632
max non-uniformity (%) |  | 3.6364
Artefacts
|  | Left monitor |  | Right monitor |  | RHS
dead pixels |  | none present |  | none present |  | 1
flicker |  | none present |  | none present

## unfors
5.7 | lx
|  |  | 0 | cd/m²
LHS
|  |  | 0.601 | cd/m² | 60 | 247741 / 219348 | 2023-03-11 | 12:59 | Units:lux & cd/m²
|  |  | 1.51 | cd/m² | 61 | 247741 / 219348 | 2023-03-11 | 12:59 | Units:lux & cd/m²
|  |  | 3.05 | cd/m² | 62 | 247741 / 219348 | 2023-03-11 | 12:59 | Units:lux & cd/m²
|  |  | 5.43 | cd/m² | 63 | 247741 / 219348 | 2023-03-11 | 12:59 | Units:lux & cd/m²
|  |  | 8.72 | cd/m² | 64 | 247741 / 219348 | 2023-03-11 | 12:59 | Units:lux & cd/m²
|  |  | 13.2 | cd/m² | 65 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 19.5 | cd/m² | 66 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 27.9 | cd/m² | 67 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 38.9 | cd/m² | 68 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 53.5 | cd/m² | 69 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 72.3 | cd/m² | 70 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 97 | cd/m² | 71 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 129 | cd/m² | 72 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 170 | cd/m² | 73 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 223 | cd/m² | 74 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 291 | cd/m² | 75 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 379 | cd/m² | 76 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
|  |  | 481 | cd/m² | 77 | 247741 / 219348 | 2023-03-11 | 13:00 | Units:lux & cd/m²
RHS
|  |  | 0.535 | cd/m² | 78 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 1.41 | cd/m² | 79 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 2.87 | cd/m² | 80 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 5.14 | cd/m² | 81 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 8.3 | cd/m² | 82 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 12.6 | cd/m² | 83 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 18.7 | cd/m² | 84 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 26.6 | cd/m² | 85 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 37 | cd/m² | 86 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 51.1 | cd/m² | 87 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 69.3 | cd/m² | 88 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 92.6 | cd/m² | 89 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 123 | cd/m² | 90 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 163 | cd/m² | 91 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 215 | cd/m² | 92 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 281 | cd/m² | 93 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 361 | cd/m² | 94 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
|  |  | 462 | cd/m² | 95 | 247741 / 219348 | 2023-03-11 | 13:02 | Units:lux & cd/m²
Uni LHS
|  |  | 2.42 | cd/m² | 96 | 247741 / 219348 | 2023-03-11 | 13:04 | Units:lux & cd/m²
|  |  | 2.39 | cd/m² | 97 | 247741 / 219348 | 2023-03-11 | 13:04 | Units:lux & cd/m²
|  |  | 2.58 | cd/m² | 98 | 247741 / 219348 | 2023-03-11 | 13:04 | Units:lux & cd/m²
|  |  | 2.39 | cd/m² | 99 | 247741 / 219348 | 2023-03-11 | 13:04 | Units:lux & cd/m²
|  |  | 2.56 | cd/m² | 100 | 247741 / 219348 | 2023-03-11 | 13:04 | Units:lux & cd/m²
|  |  | 200 | cd/m² | 101 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 200 | cd/m² | 102 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 196 | cd/m² | 103 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 199 | cd/m² | 104 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 197 | cd/m² | 105 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
Uni RHS 10?80 |  |  | 2.29 | cd/m² | 106 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 2.29 | cd/m² | 107 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 2.53 | cd/m² | 108 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 2.32 | cd/m² | 109 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 2.3 | cd/m² | 110 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 2.3 | cd/m² | 111 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 190 | cd/m² | 112 | 247741 / 219348 | 2023-03-11 | 13:05 | Units:lux & cd/m²
|  |  | 196 | cd/m² | 113 | 247741 / 219348 | 2023-03-11 | 13:06 | Units:lux & cd/m²
|  |  | 192 | cd/m² | 114 | 247741 / 219348 | 2023-03-11 | 13:06 | Units:lux & cd/m²
|  |  | 195 | cd/m² | 115 | 247741 / 219348 | 2023-03-11 | 13:06 | Units:lux & cd/m²
|  |  | 189 | cd/m² | 116 | 247741 / 219348 | 2023-03-11 | 13:06 | Units:lux & cd/m²
LHS TG-18
|  | 1 | 474 | cd/m² | 117 | 247741 / 219348 | 13/07/2022 | 15:52 | Units:lux & cd/m²
|  | 2 | 472 | cd/m² | 156 | 247741 / 219348 | 13/07/2022 | 15:59 | Units:lux & cd/m²
|  | 3 | 470 | cd/m² | 157 | 247741 / 219348 | 13/07/2022 | 15:59 | Units:lux & cd/m²
|  | 4 | 474 | cd/m² | 158 | 247741 / 219348 | 13/07/2022 | 15:59 | Units:lux & cd/m²
|  | 5 | 477 | cd/m² | 159 | 247741 / 219348 | 13/07/2022 | 15:59 | Units:lux & cd/m²
|  | 6 | 478 | cd/m² | 160 | 247741 / 219348 | 13/07/2022 | 15:59 | Units:lux & cd/m²
|  | 7 | 479 | cd/m² | 161 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²
|  | 8 | 477 | cd/m² | 162 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²
|  | 9 | 474 | cd/m² | 163 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²
|  | 10 | 472 | cd/m² | 164 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²
|  | 11 | 477 | cd/m² | 165 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²
|  | 12 | 479 | cd/m² | 166 | 247741 / 219348 | 13/07/2022 | 16:00 | Units:lux & cd/m²